## Notification Sheet
Technical Measurement Issue Notification
Significant Measurement Error ID |  | SC006
Query Issue ID (Allocated by Joint Office) |  | 001
Submitting Party | Organisation: | BRITISH GAS
| Contact Name: | GRAHAM WOOD
| Business Number: | 07979 567686
| email address | graham.wood@britishgas.co.uk
Fax Number:
Technical Measurement Issue Description | Following the presentation made on 16th July 2012, we have a number of questions, detailed below for Ben Kirkham only.
Anticipated Impact on Critical Data Used in Evaluation of Measurement Error Magnitude | Essential questions to provide Shippers with further detail and background to work done to date and future work to be undertaken by the ITE.
Submission Date |  | 24th August 2012
1. A claim was made that flow before and after could not be determined, maybe this needs to be looked at a longer time period. Flow was 20 mmscf/d before plate returned to its correct position then 40 mmscf/d afterward, how much of the flow is due to transient demand at the time?
2. No spread or uncertainty has been stated stated for derivation of error, we would be interested in your view on this.
3. We are mindful that there is the potential that additional site tests may be required, what is your view on this? We are clear that such a requirement should not be discounted on grounds of cost or inconvenience, due to the value of the error. If further site tests are required then this should be clearly recommended by the ITE.
4. How were the counter readings verified? This is a critical aspect to establish the exact location of the plate.
5. The name plate included a hard to read counter reference and the counters first and last digits were also “clouded” / easily mis-read. So, while in on a scale of 1-100 how confident are you that the “perceived” plate position while in correctly positioned, was what you have presumed: Case 1? Case2?
6. The CFD model is questionable in terms of how it is used. We would expect it to be very close to the standard ISO 5167 calculation as this is well defined. However when the plate is not in the correct position it is outside of ISO 5167 so there are no rules or guidance for the calculation. It may therefore need another independent calculation as a check for the first once?
7. Is the CFD expert subject to ISO 9001, to include peer review and who was doing that review?
8. As the CFD results for the “test” point have been unsuccessful so far, how many attempts are you going to give the expert, before this is seen as a “lucky” answer and this line of enquiry no longer progressed?
9. Do the test flows and line conditions match that of the meter run at the time, especially in a cold winter (2010)?
10. As the pressure loss is far less than expected for the flow rate how is the upstream temperature determined? With ISO 5167 this is clearly defined, but with the plate not in place is can only be assumed, this has an impact as the lack of valid temperature correction will affect the calculated gas density as the upstream temperature has to be determined from an known model with an inconsistent flow profile in the meter run.
11. The orifice plates removed may also provide evidence but we have not yet seen any photos, are there any that can be shared?

## Extract for JO Database
Measurement Error ID | Query Issue ID | Organisation | Contact Name | Business Number | email address | Fax Number | Technical Measurement Issue Description | Anticipated Impact on Critical Data Used in Evaluation of Measurement Error Magnitude | Submission Date
SC006 | SC006 | BRITISH GAS | GRAHAM WOOD | 07979 567686 | 0 | 0 | Following the presentation made on 16th July 2012, we have a number of questions, detailed below for Ben Kirkham only. | Essential questions to provide Shippers with further detail and background to work done to date and future work to be undertaken by the ITE. | 24th August 2012